## Slide 1
Carpediem™
Cardio-Renal Pediatric Dialysis
Emergency System
  Machine overview, components, disposables and consumables
  resource
  June 2020

## Slide 2: 2012 first in the world
Prof. Claudio Ronco Vincenza Italy

## Slide 3: Carpediem™ Pediatric Dialysis system
Indicated in acute kidney injury or fluid overloaded patients requiring hemodialysis or hemofiltration therapy
The Carpediem™ system is capable of performing ultrafiltration, hemofiltration and hemodialysis
Intended to provide continuous renal replacement therapy (CRRT) for pediatric patients weighing between 2.5-10 kilograms (5.5 lbs–22 lbs)
Special miniaturized components ensure high accuracy in weight loss and pump flows
Dedicated Single-use kits support low priming volumes
Overview
Overview | Components | Disposables & consumables
Notes: .

## Slide 4
SCUF: Slow Continuous Ultrafiltration
CVVH: Continuous Veno-Venous Hemofiltration
      Pre dilution (pre dialyzer)
      Post dilution (post dialyzer)
CVVHD: Continuous Veno-Venous Hemodialysis
TREATMENT OPTIONS
Carpediem™ Pediatric Dialysis
system
Both preassembled device models may be used to deliver SCUF, CVVH and CVVHD
Overview | Components | Disposables & consumables

## Slide 5: components

## Slide 6: Carpediem™ system components
Back View and cords
Base plate
Potential equalizer
point
Protection ground connector
Power supply plug
Mains ON (‘I’)/
OFF (‘O’) switch
Warmer holder
Labels
Overview | Components | Disposables & consumables

## Slide 7: Carpediem™ system components
Cords and power
Note: When using a CVC access for treatment (with a tip in the right atrium), the use of the equipotential cable is mandatory.
The cord must be connected to the machine and then connected/plugged into a grounded electrical hospital socket
Equipotential cord
Overview | Components | Disposables & consumables

## Slide 8
WARNING: Battery should be recharged ONCE A MONTH for at least 8 hours . Refer to Carpediem™ Operators Manual for further information
Battery
Stores data for 2 minutes
Not designed to run a therapy
Battery check/test recommend prior to treatment:
  If NO alarm, recharge battery by plugging into power socket
  If alarm 21 BATTERY DEAD presents, RECHARGE before use
Remove cord
Carpediem™ system components
Overview | Components | Disposables & consumables

## Slide 9
05-2019 CARPEDIEM REV. 01 AUTHOR: SILVIA MANFREDINI
Front view
Effluent pump
Blood pump
Infusion/dialysis (INF/DIA) pump
Arterial pressure transducer (Dome)
Venous pressure transducer (Dome)
Pre-filter pressure transducer (Dome)
Clip for infusion/dialysis pump segment
Clip for venous line
Clip for infusion/dialysis line if a warmer is used
Clip for effluent pump segment
Filter holder
Venous drip chamber holder
Air sensor
Venous electroclamp
Blood leakage detector (BLD)
Heparin pump
Effluent (EFF) scale
Infusion/dialysis (INF/DIA) scale
Control panel
Warning light
IV pole holder
Warmer holder
Carpediem™ system components
Overview | Components | Disposables & consumables

## Slide 10: Carpediem™ system components
Inside the Carpediem machine
Overview | Components | Disposables & consumables
All parts designed to:
Increase miniaturization
Reduce priming volume
Reduce circuit pressure

## Slide 11: Carpediem™ system Components
High performance pumps
3 micronized peristaltic pumps:
    Blood pump
    Effluent pump
    Infusion/dialysate pump
Each pump is equipped with 3 rollers to maximize delivery precision
Blood pump
Infusion/
Dialysis
pump
Effluent pump
Overview | Components | Disposables & consumables
Peristaltic pump utilizes the movement of the stator to increase miniaturization

## Slide 12
Efficiency pumps
Miniaturization considerations
  Occlusion
  Pressure peaks
  Filling volume
Carpediem™ system Components
VS
Overview | Components | Disposables & consumables

## Slide 13
Pressure Sensor transducers
3 unique pressure transducers (Dome):
  Arterial
  Prefilter
  Venous
Protective dome covers prevent damage to the pressure transducer domes
Locking mechanism holds tubing in place
Carpediem™ system components
Overview | Components | Disposables & consumables
Each pressure transducer is linked to a pump
Pressure dome-back
Pressure dome- front
Membrane

## Slide 14
Bolus: 0.1 to 3 mL – resolution (increments) of 0.1 mL
Continuous: 0.4 to 20 mL/h – resolution (increments) of 0.1 mL
Single-use 10 mL syringe is provided in the Single-use Kit
HEPARIN OPPERATING RANGES and SYRINGE
Note:
ONLY the validated 10 mL syringe provided in the Single-use kit can be used with the Carpediem™ system
ONLY Heparin has been validated for use with Carpediem™ system
Carpediem™ system Components
Overview | Components | Disposables & consumables

## Slide 15: Carpediem™ system components
Safety
The system limits the convective rate to 20% of the blood flow recognizing the type of hemofilter thus preventing excessive hemoconcentration
The system recognizes the type of preassembled device and delivers the appropriate volume during blood restitution
Weight loss accuracy is limited to + 30g per 24 hrs.
Overview | Components | Disposables & consumables

## Slide 16: Carpediem™ system components
scales
Fluid balance is controlled and managed by the infusion and effluent scales with a resolution of 1 gram
The maximum scale/weight capacity 2.4 Kg
Effluent scale
Infusion/dialysis scale
Overview | Components | Disposables & consumables

## Slide 17
Operating ranges
Weight loss: 0 to 2000 g
Weight loss accuracy: +30 grams in 24hrs
Max scale capacity: 2400 g
Treatment time: maximum treatment time
of 24 hrs; increments of 30 minutes
Carpediem™ Components
Overview | Components | Disposables & consumables
Effluent (EFF)
scale
Effluent pump
Blood pump
Infusion/dialysis (INF/DIA) scale
Infusion/dialysis (INF/DIA) pump

## Slide 18
Operating ranges
Blood pump flow rate: from 2 to 50 mL/min, resolution (increments) of 1 mL/min
Infusion/dialysate pump flow rate: 0 to 600 mL/h, resolution (increments) of 1 mL/min
Effluent pump flow rate: 0 to 900 mL/h, resolution (increments) of 1 mL/min
Carpediem™ system Components
Overview | Components | Disposables & consumables
CVVH |  | CVVHD | 
HCD 015 Filter | From 0 to 4mL/min (0 to 240 mL/hr.) | HCD 015 Filter | From 0 to 10mL/min (0 to 600 mL/hr.)
HCD 025 Filter | From 0 to 10mL/min (0 to 600 mL/hr.) | HCD 025 Filter | From 0 to 10mL/min (0 to 600 mL/hr.)
CVVH |  | CVVHD | 
HCD 015 Filter | From 0 to 4mL/min (0 to 240 mL/hr.) | HCD 015 Filter | From 0 to 14mL/min (0 to 840 mL/hr.)
HCD 025 Filter | From 0 to 10mL/min (0 to 600 mL/hr.) | HCD 025 Filter | From 0 to 15mL/min (0 to 900 mL/hr.)
NOTE: Effluent pump flow is limited to maximum 20% of the blood flow in (CVVH)
NOTE: Weight loss flow is limited to maximum 20% of the blood flow

## Slide 19
A combination of simultaneous audible and visual cues communicate the state of the machine
The warning light will change color to communicate machine status or the priority of an alarm
A buzz will be heard during an alarm state
The message center will communicate information to the user before, during and after treatment
The help/information key ‘?’ provides information or direction on possible alarm causes as well as potential solution(s)
Machine Status & Alarms communication
Warning light
Carpediem™ System components
Machine interface: Priming
Overview | Components | Disposables & consumables

## Slide 20
YELLOW and RED are used to define warning and alarm priorities
The blood pump may stop, slow or continue to operate depending on the alarm
Note: If more than one alarm state exists, the highest priority alarm will present first. Once the higher priority alarm has been resolved, the lower priority alarm will be shown
Alarm Color and priority
Message Color | Alarms identified | Results
Yellow 0.8 Hz | MEDIUM PRIORITY | Blood pump continues to operate; Tx running despite alarm
Red 1.6 Hz | HIGH PRIORITY | Blood pump stops or slows; Rapid intervention by operator required to correct issue or prevent coagulation of the blood circuit
Green light = NO alarm
Color and sound used to identify status conditions
Message class | Visual Signal | Audible Signal | Description
No Alarm | Green Continuous | Absent | Indicates that the equipment is in normal operating condition
Type I | Green Flashing | Absent | Indicates that the equipment is running a test (or an override) and does not require intervention by the operator
Type II | Green Flashing | Presents with a pulse every 15 sec | Indicates that the equipment has suspended the treatment and requires intervention by the operator
Overview | Components | Disposables & consumables
Carpediem™ System components

## Slide 21
Help/information
button
Alarm silence/reset
button
Adjustment and selection Knob
Start/Stop
button
CONTROL PANEL
Carpediem™ system components
The control panel is used to set parameters and manage the operator interface
Overview | Components | Disposables & consumables

## Slide 22
Once initial testing completed, the operator MUST choose one of three available options:
Control Panel
OPTION 1 | OPTION 2 | OPTION 3
CVVH | CVVHD | Technical Parameters
Language, date,
time, display contrast
Note: a SCUF treatment can be performed selecting CVVH and then setting infusion volume to zero (0)
Note: Technical Parameters should be chosen the first time the machine is set up or language changed
Carpediem™ System Components
Mode of therapy
Common Procedures | Adjunct Procedures

## Slide 23
Operator Interface
Modifiable fields have white backgrounds; fixed fields are grey
To modify a parameter, turn the knob on the display panel until the desired field background turns yellow
Press the knob to set the field highlighted
Turn the selection knob to the right or left to change the value
Press the knob once more to confirm/save the programming; background will turn blue following confirmation
Note: the field background turns blue when parameter confirmed
Selected field
Confirmed field modification
Modifiable fields
Adjustment and selection knob
English
dd
hh
Carpediem™ System Components
English
Overview | Components | Disposables & consumables

## Slide 24: Possible priming alarm
If the alarms 15 EFFLUENT WEIGHT NOT ZERO occurs during the system preparation when a bag has been hung on EFF scale prior to installing the lines.
The scales must not have any weight on them on until the treatment mode has been selected
Remove the bags, reset or turn off and on machine using white ‘on/off’ key
Scale weights do not read zero
Alarms Management | Common Alarms | Emergency Procedures | Trouble Shooting
15 EFFLUENT WEIGHT NOT ZERO

## Slide 25: Possible priming alarm
If the alarm 16 INF/DIA WEIGHT NOT ZERO occurs during the system preparation, infusion bags have been hung on the INF/DIA scale prior to installing the lines.
The scales must not have any weight on them on until the treatment mode has been selected
Remove the bags, reset or turn off and on machine using white ‘on/off’ key
Scale weights do not read zero
Alarms Management | Common Alarms | Emergency Procedures | Trouble Shooting
16 INF/DIA WEIGHT NOT ZERO

## Slide 26
No one specific alarm
Check circuit installation
Check that the bags are properly positioned on their appropriate scales
Check that the return line (venous) is properly positioned in the air sensor
Check that the ultrafiltration line (effluent) is properly positioned in the BLD sensor
Check for air bubbles in the section of line monitored by the BLD sensor
Check pressure sensor functionality by slightly closing (manual clamping) sections of the circuit monitored by the sensors
Check which test did not pass:
    Protection relay test
    Pressurization test
Should the priming phase not complete, select End Priming and contact Technical services
Possible priming alarm
Incomplete priming procedure
Alarms Management | Common Alarms | Emergency Procedures | Trouble Shooting

## Slide 27: Disposables & consumables

## Slide 28: Carpediem™ system Disposables and Consumables
Carpediem™ Single-use kit
Carpediem™ Single-use Kit
Includes:
  Blood lines with attached dialyzer(filter)
    HCD 015 model
    or
    HCD 025 model
  3 X 3L sterile waste bags
  1 X 10 mL Syringe
  Single-use Kit IFU
  Preassembled Device IFU
Overview | Components | Disposables & consumables
10 mL syringe
Dialyzer (filter)
Waste bags
IFUs
Carpediem™ Single-Use
kit packaging
‘PRE’, ‘POST’ and ‘DIA’ labels
labels
label
Carpediem™ Single-Use kit
Venous chamber
Pressure dome
Pressure dome
Pressure dome

## Slide 29
Arterial line sampling port
Arterial line Luer-lock
Effluent line Luer-lock
Effluent line
Effluent pump segment
Venous pressure dome
Infusion/dialysis line Luer-lock
Infusion/dialysis pump segment
Venous chamber
Venous return line sampling port
Venous line Luer-lock
Venous return line
Blood pump segment
Arterial pressure dome
Pre-filter pressure dome
Filter
Heparin pump line Luer-lock
Heparin line
PREASSEMBLED BLOOD LINE SCHEMATIC
Carpediem™ system Disposables and Consumables
Overview | Components | Disposables & consumables
PRE
DIA
POST
Arterial supply line

## Slide 30
Disposable model Options
Preassembled device description | Filter Membrane material | Filter Surface (m2) | Preassembled Device Priming volume (mL) | Maximum Eff pump flow (mL/min) | Maximum EFF pump flow (mL/hr.) | Sterilization
015 CVVH/CVVHD | Polyethersulfone | 0.16 | 32 | 4 (CVVH) 14 (CVVHD) | 240 (CVVH) 840 (CVVHD | EtO
025 CVVH/CVVHD | Polyethersulfone | 0.29 | 41 | 10 (CVVH) 15 (CVVHD) | 600 (CVVH) 900 (CVVHD) | EtO
Overview | Components | Disposables & consumables
Carpediem™ system Disposables and Consumables

## Slide 31
Consumable- dialysate fluid bag
Sterile solution
Two chamber bag
K 2.5 mmol/L
HCO3 of 32mmol/L
Na 140 mmol/L
Ca 1.5 mmol/L
Carpediem™ System Disposables and Consumables
Clear outer
packaging
Double chambered bag
Overview | Components | Disposables & consumables

## Slide 32: Carpediem™ system Disposables and Consumables
Consumable- dialysate fluid bag
Clear outer packaging
Double chambered bag
Overview | Components | Disposables & consumables
Breakable cone
18-month shelf life
Not for injection
Intact cone
Broken cone

## Slide 33
Use of a warmer is optional
Set up prior to priming
Optional bIEGLER
warmer
An equipopotential cord is required if the patient access is a CVC
Connection to the preassembled device is made at the point of the infusion/dialysis line- not blood line
Requires dedicated disposable tubing
Connects to the pole on the machine base plate
Carpediem™ system Components
Overview | Components | Disposables & consumables

## Slide 34: Resources and Support

## Slide 35: Clinical resources
Clinical procedure and trouble shooting resource
Carpediem™ machine Operator’s Manual
Access to a clinician

## Slide 36: Emergency phone numbers
CLINICAL support: 24 hours a day, 7 days a week
CLINICAL Support Phone Number 24/7 Accessibility
1-800-255-6774
Select option 5
Select option 2
1-800-255-6774
Select option 5
Select option 1
Calls received outside of office hours
directed to voicemail
A technical representative will return the
call within 2 hours the next business day
TECHNICAL Support Phone Number (Office hrs. 6AM-4PM PST)
TECHNICAL support: Monday-Friday 6AM-4PM PST

## Slide 37: ADDITIONAL Contacts
CLINICAL Support: 24/7 Monday-Sunday
(All clinical support questions)
  Telephone: +1-800-255-6774; follow the prompts and selection options
  e-mail: rs.rcsclinicalsupport@medtronic.com
TECHNICAL Support: Office hours are Monday-Friday from 6AM-4PM PST
(Repair requests, service status, other non-emergency technical questions)
  Telephone: +1-800-255-6774; follow the prompts and selection options
  Fax: +1-760-603-5816
  e-mail: rs.serviceexperience@medtronic.com
CUSTOMER Service: Office hours Monday-Friday 8 AM- 6:30 PM EST
(Place an order, order status, tracking shipments)
  Telephone : +1-800-962-9888
  e-mail: rs.covidiencustomerservice@medtronic.com
Non-Emergency Contact Information

## Slide 38
  Bellco S.r.l. (2020). US Carpediem™ Operator’s Manual (rev 10). SW 1.5-0 IB5332941ENG
  Bellco S.r.l. (2019). CVVH and CVVHD Short Manual Carpediem (rev. 04). SW 1.5-0 IBSHM3001. Taken from Operator’s Manual IB5332941ENG Rev. 02 ed. 04/19.
  Bellco S.r.l.(2019). Single Use Kit Information For Use (rev.00). IB012002800
  Bellco S.r.l. (2020). Preassembled Device Information For Use. IB012002700
  Haemopharm Biofluids s.r.l.(2015). Hbiofluids Instruction For Use. Rev.01.
© 2020 Medtronic. All rights reserved. Medtronic, Medtronic logo and Further, Together are trademarks of Medtronic.
All other brands are trademarks of a Medtronic company. 06/2020 – US-RC-2000017
References